# Purpose

The purpose of this document is to define the device classification of the <device name> device.

# Applicable Documents

  - Medical Device Regulation (EU) 2017/745 (MDR)

  - MDCG 2021-24, Guidance on classification of medical devices, issued in October 2021

# Definitions

Medical device – any instrument, apparatus, appliance, software, implant, reagent, material or other article intended by the manufacturer to be used, alone or in combination, for human beings for one or more of the following specific medical purposes:

- diagnosis, prevention, monitoring, prediction, prognosis, treatment or alleviation of disease,

- diagnosis, monitoring, treatment, alleviation of, or compensation for, an injury or disability,

- investigation, replacement or modification of the anatomy or of a physiological or pathological process or state,

- providing information by means of in vitro examination of specimens derived from the human body, including organ, blood and tissue donations,

and which does not achieve its principal intended action by pharmacological, immunological or metabolic means, in or on the human body, but which may be assisted in its function by such means.

The following products shall also be deemed to be medical devices:

- devices for the control or support of conception;

- products specifically intended for the cleaning, disinfection or sterilisation of devices as referred to in Article 1(4) of MDR 2017/745 and of those referred to in the first paragraph of this point.

Duration of use:

  - Transient - normally intended for continuous use for less than 60 minutes.

  - Short term - normally intended for continuous use for between 60 minutes and 30 days.

  - Long term - normally intended for continuous use for more than 30 days.

Invasive and active devices:

  - Active device - any device, the operation of which depends on a source of energy other than that generated by the human body for that purpose, or by gravity, and which acts by changing the density of or converting that energy. Devices intended to transmit energy, substances or other elements between an active device and the patient, without any significant change, shall not be deemed to be active devices. Software shall also be deemed to be an active device.

  - Active device intended for diagnosis and monitoring - any active device used, whether alone or in combination with other devices, to supply information for detecting, diagnosing, monitoring or treating physiological conditions, states of health, illnesses or congenital deformities.

  - Active therapeutic device: any active device used, whether alone or in combination with other devices, to support, modify, replace or restore biological functions or structures with a view to treatment or alleviation of an illness, injury or disability.

  - Body orifice - any natural opening in the body, as well as the external surface of the eyeball, or any permanent artificial opening, such as a stoma.

  - Central circulatory system - the following blood vessels: arteriae pulmonales, aorta ascendens, arcus aortae, aorta descendens to the bifurcatio aortae, arteriae coronariae, arteria carotis communis, arteria carotis externa, arteria carotis interna, arteriae cerebrales, truncus brachiocephalicus, venae cordis, venae pulmonales, vena cava superior and vena cava inferior.

  - Central nervous system - the brain, meninges and spinal cord.

  - Injured skin or mucous membrane - an area of skin or a mucous membrane presenting a pathological change or change following disease or a wound.

  - Invasive device - any device which, in whole or in part, penetrates inside the body, either through a body orifice or through the surface of the body.

  - Surgically invasive device -

(a) an invasive device which penetrates inside the body through the surface of the body, including through mucous membranes of body orifices with the aid or in the context of a surgical operation; and

(b) a device which produces penetration other than through a body orifice.

  - Reusable surgical instrument - an instrument intended for surgical use in cutting, drilling, sawing, scratching, scraping, clamping, retracting, clipping or similar procedures, without a connection to an active device and which is intended by the manufacturer to be reused after appropriate procedures such as cleaning, disinfection and sterilisation have been carried out.

# EU Device Classification

The <Device name> is intended for <intended use>. <Insert short description>

The <Device name> is a medical device because it is intended to <rationale here>.

The <Device name> is classified according to Annex VIII of the Medical Device Regulation (EU) 2017/745 (MDR) as Class X according to Rule X.

# Implementing Rules

All rules detailed in Annex VIII of the 2017/745 Medical Device Regulation are reviewed in Appendix 1 of this document.

# EU Conformity Assessment Route

Annex IX (chapter X) of the Medical Device Regulation (EU) 2017/745 (MDR) and EN ISO 13485:2016.

# Responsibilities

The Personal Responsible for Regulatory Compliance (PRRC) is responsible for ensuring that the classification of the device is properly defined.

| Appendix I - Classification rules of the MDR Annex VIII |  |  |  |
| --- | --- | --- | --- |
| Rule # | Rule description | Class | Classification (applicability and rational) |
| Non-Invasive Devices |  |  |  |
| 1 | All non-invasive devices are classified as class I, unless one of the rules set out hereinafter applies. | I |  |
| 2 | All non-invasive devices intended for channelling or storing blood, body liquids, cells or tissues, liquids or gases for the purpose of eventual infusion, administration or introduction into the body are classified as class IIa: - if they may be connected to a class IIa, class IIb or class III active device; or - if they are intended for use for channelling or storing blood or other body liquids or for storing organs, parts of organs or body cells and tissues, except for blood bags; | IIa |  |
|  | blood bags are classified as class IIb. | IIb |  |
|  | In all other cases, such devices are classified as class I. | I |  |
| 3 | All non-invasive devices intended for modifying the biological or chemical composition of human tissues or cells, blood, other body liquids or other liquids intended for implantation or administration into the body are classified as class IIb, unless the treatment for which the device is used consists of filtration, centrifugation or exchanges of gas, heat, in which case they are classified as class IIa. | IIa/IIb |  |
|  | All non-invasive devices consisting of a substance or a mixture of substances intended to be used in vitro in direct contact with human cells, tissues or organs taken from the human body or used in vitro with human embryos before their implantation or administration into the body are classified as class III. | III |  |
| 4 | All non-invasive devices which come into contact with injured skin or mucous membrane are classified as: |  |  |
|  | - class I if they are intended to be used as a mechanical barrier, for compression or for absorption of exudates; | I |  |
|  | - class IIb if they are intended to be used principally for injuries to skin which have breached the dermis or mucous membrane and can only heal by secondary intent; | IIb |  |
|  | - class IIa if they are principally intended to manage the micro-environment of injured skin or mucous membrane; and | IIa |  |
|  | - class IIa in all other cases. | IIa |  |
|  | This rule applies also to the invasive devices that come into contact with injured mucous membrane. |  |  |
| Invasive Devices |  |  |  |
| 5 | All invasive devices with respect to body orifices, other than surgically invasive devices, which are not intended for connection to an active device or which are intended for connection to a class I active device are classified as: |  |  |
|  | - class I if they are intended for transient use; | I |  |
|  | - class IIa if they are intended for short-term use, except if they are used in the oral cavity as far as the pharynx, in an ear canal up to the ear drum or in the nasal cavity, in which case they are classified as class I; and | IIa |  |
|  | - class IIb if they are intended for long-term use, except if they are used in the oral cavity as far as the pharynx, in an ear canal up to the ear drum or in the nasal cavity and are not liable to be absorbed by the mucous membrane, in which case they are classified as class IIa. | IIb |  |
|  | All invasive devices with respect to body orifices, other than surgically invasive devices, intended for connection to a class IIa, class IIb or class III active device, are classified as class IIa. | IIa |  |
| 6 | All surgically invasive devices intended for transient use are classified as class IIa unless they: | IIa |  |
|  | - are intended specifically to control, diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with those parts of the body, in which case they are classified as class III; | III |  |
|  | - are reusable surgical instruments, in which case they are classified as class I; | I |  |
|  | - are intended specifically for use in direct contact with the heart or central circulatory system or the central nervous system, in which case they are classified as class III; | III |  |
|  | - are intended to supply energy in the form of ionising radiation in which case they are classified as class IIb; | IIb |  |
|  | - have a biological effect or are wholly or mainly absorbed in which case they are classified as class IIb; or | IIb |  |
|  | - are intended to administer medicinal products by means of a delivery system, if such administration of a medicinal product is done in a manner that is potentially hazardous taking account of the mode of application, in which case they are classified as class IIb. | IIb |  |
| 7 | All surgically invasive devices intended for short-term use are classified as class IIa unless they: | IIa |  |
|  | - are intended specifically to control, diagnose, monitor or correct a defect of the heart or of the central circulatory system through direct contact with those parts of the body, in which case they are classified as class III; | III |  |
|  | - are intended specifically for use in direct contact with the heart or central circulatory system or the central nervous system, in which case they are classified as class III; | III |  |
|  | - are intended to supply energy in the form of ionizing radiation in which case they are classified as class IIb; | IIb |  |
|  | - have a biological effect or are wholly or mainly absorbed in which case they are classified as class III; | III |  |
|  | - are intended to undergo chemical change in the body in which case they are classified as class IIb, except if the devices are placed in the teeth; or | IIb |  |
|  | - are intended to administer medicines, in which case they are classified as class IIb. | IIb |  |
| 8 | All implantable devices and long-term surgically invasive devices are classified as class IIb unless they: | IIb |  |
|  | - are intended to be placed in the teeth, in which case they are classified as class IIa; | IIa |  |
|  | - are intended to be used in direct contact with the heart, the central circulatory system or the central nervous system, in which case they are classified as class III; | III |  |
|  | - have a biological effect or are wholly or mainly absorbed, in which case they are classified as class III; | III |  |
|  | - are intended to undergo chemical change in the body in which case they are classified as class III, except if the devices are placed in the teeth; | III |  |
|  | - are intended to administer medicinal products, in which case they are classified as class III; | III |  |
|  | - are active implantable devices or their accessories, in which cases they are classified as class III; | III |  |
|  | - are breast implants or surgical meshes, in which cases they are classified as class III; | III |  |
|  | - are total or partial joint replacements, in which case they are classified as class III, with the exception of ancillary components such as screws, wedges, plates and instruments; or | III |  |
|  | - are spinal disc replacement implants or are implantable devices that come into contact with the spinal column, in which case they are classified as class III with the exception of components such as screws, wedges, plates and instruments. | III |  |
| Active Devices |  |  |  |
| 9 | All active therapeutic devices intended to administer or exchange energy are classified as class IIa unless their characteristics are such that they may administer energy to or exchange energy with the human body in a potentially hazardous way, taking account of the nature, the density and site of application of the energy, in which case they are classified as class IIb. | IIa/IIb |  |
|  | All active devices intended to control or monitor the performance of active therapeutic class IIb devices, or intended directly to influence the performance of such devices are classified as class IIb. | IIb |  |
|  | All active devices intended to emit ionizing radiation for therapeutic purposes, including devices which control or monitor such devices, or which directly influence their performance, are classified as class IIb. | IIb |  |
|  | All active devices that are intended for controlling, monitoring or directly influencing the performance of active implantable devices are classified as class III. | III |  |
| 10 | Active devices intended for diagnosis and monitoring are classified as class IIa: - if they are intended to supply energy which will be absorbed by the human body, except for devices intended to illuminate the patient's body, in the visible spectrum, in which case they are classified as class I; if they are intended to image in vivo distribution of radiopharmaceuticals; or - if they are intended to allow direct diagnosis or monitoring of vital physiological processes, unless they are specifically intended for monitoring of vital physiological parameters and the nature of variations of those parameters is such that it could result in immediate danger to the patient, for instance variations in cardiac performance, respiration, activity of the central nervous system, or they are intended for diagnosis in clinical situations where the patient is in immediate danger, in which cases they are classified as class IIb. | I/IIa/IIb |  |
|  | Active devices intended to emit ionizing radiation and intended for diagnostic or therapeutic radiology, including interventional radiology devices and devices which control or monitor such devices, or which directly influence their performance, are classified as class IIb. | IIb |  |
| 11 | Software intended to provide information which is used to take decisions with diagnosis or therapeutic purposes is classified as class IIa, except if such decisions have an impact that may cause: | IIa |  |
|  | - death or an irreversible deterioration of a person's state of health, in which case it is in class III; or | III |  |
|  | - a serious deterioration of a person's state of health or a surgical intervention, in which case it is classified as class IIb. | IIb |  |
|  | Software intended to monitor physiological processes is classified as class IIa, except if it is intended for monitoring of vital physiological parameters, where the nature of variations of those parameters is such that it could result in immediate danger to the patient, in which case it is classified as class IIb. | IIa /IIb |  |
| 12 | All active devices intended to administer and/or remove medicinal products, body liquids or other substances to or from the body are classified as class IIa, unless this is done in a manner that is potentially hazardous, taking account of the nature of the substances involved, of the part of the body concerned and of the mode of application in which case they are classified as class IIb. | IIa/IIb |  |
| 13 | All other active devices are classified as class I. | I |  |
| Special Rules |  |  |  |
| 14 | All devices incorporating, as an integral part, a substance which, if used separately, can be considered to be a medicinal product, as defined in point 2 of Article 1 of Directive 2001/83/EC, including a medicinal product derived from human blood or human plasma, as defined in point 10 of Article 1 of that Directive, and that has an action ancillary to that of the devices, are classified as class III. | III |  |
| 15 | All devices used for contraception or prevention of the transmission of sexually transmitted diseases are classified as class IIb, unless they are implantable or long term invasive devices, in which case they are classified as class III. |  |  |
| 16 | All devices intended specifically to be used for disinfecting, cleaning, rinsing or, where appropriate, hydrating contact lenses are classified as class IIb. | IIb |  |
|  | All devices intended specifically to be used for disinfecting or sterilising medical devices are classified as class IIa, unless they are disinfecting solutions or washer-disinfectors intended specifically to be used for disinfecting invasive devices, as the end point of processing, in which case they are classified as class IIb. This rule does not apply to devices that are intended to clean devices other than contact lenses by means of physical action only. | IIa/IIb |  |
| 17 | Devices specifically intended for recording of diagnostic images generated by X-ray radiation are classified as class IIa. | IIa |  |
| 18 | All devices manufactured utilising tissues or cells of human or animal origin, or their derivatives, which are non-viable or rendered non-viable, are classified as class III, unless such devices are manufactured utilising tissues or cells of animal origin, or their derivatives, which are non-viable or rendered non-viable and are devices intended to come into contact with intact skin only. | III |  |
| 19 | All devices incorporating or consisting of nanomaterial are classified as: |  |  |
|  | - class III if they present a high or medium potential for internal exposure; | III |  |
|  | - class IIb if they present a low potential for internal exposure; and | IIb |  |
|  | - class IIa if they present a negligible potential for internal exposure. | IIa |  |
| 20 | All invasive devices with respect to body orifices, other than surgically invasive devices, which are intended to administer medicinal products by inhalation are classified as class IIa, unless their mode of action has an essential impact on the efficacy and safety of the administered medicinal product or they are intended to treat life-threatening conditions, in which case they are classified as class IIb. | IIa/IIb |  |
| 21 | Devices that are composed of substances or of combinations of substances that are intended to be introduced into the human body via a body orifice or applied to the skin and that are absorbed by or locally dispersed in the human body are classified as: |  |  |
|  | - class III if they, or their products of metabolism, are systemically absorbed by the human body in order to achieve the intended purpose; | III |  |
|  | - class III if they achieve their intended purpose in the stomach or lower gastrointestinal tract and they, or their products of metabolism, are systemically absorbed by the human body; | III |  |
|  | - class IIa if they are applied to the skin or if they are applied in the nasal or oral cavity as far as the pharynx, and achieve their intended purpose on those cavities; and | IIa |  |
|  | - class IIb in all other cases. | IIb |  |
| 22 | Active therapeutic devices with an integrated or incorporated diagnostic function which significantly determines the patient management by the device, such as closed loop systems or automated external defibrillators, are classified as class III. | III |  |